## Sheet1
Date EUA Original Issued | Test Developer | Name of Test | Commercial | Authorized Setting | Target Gene(s) for SARS-CoV-2 | Limit of Detection (LoD) | Product LoD (NDU/ml)* | Time to results** | Specimen Indicated for Testing | Recommended RT-PCR Instrument (analysis software) or Detection Instrument | Criteria for Positive Interpretation of SARS-CoV-2 | Remarks | Reference
Real-time RT-PCR
2020-01-10 | UMass Memorial Medical Center | UMass Molecular Virology Laboratory 2019-nCoV rRT-PCR Dx Panel | No | H | N1, N2 | 8 copies/µL | N/A | N/A | Nasal, mid-turbinate nasal, NP and OP swab, NP aspirate, and BAL specimens | BioRAD CFX96 Touch Real-Time PCR Detection System | Positive for both of the target genes |  | https://www.fda.gov/media/142692/download
2020-01-10 | Tempus Labs, Inc. | iC SARS-CoV2 Test | No | H | Orf1ab, S, N | 250 copies/mL | N/A | 2-4 hours | Upper respiratory tract specimens (including NP, anterior nares (AN or nasal), midturbinate nasal, and OP swab specimens) | Applied Biosystem QuantStudio 7 Flex 384 instrument (software version 1.3) | Positive for two or more of the target genes |  | https://www.fda.gov/media/142596/download
09/30/2020 | Aeon Global Health | Aeon Global Health SARS-CoV-2 Assay | No | H | Orf1ab, S, N | 0.25 copies/µL | N/A | 2-4 hours | Nasal, mid-turbinate, NP, and OP swab specimens, NP wash/aspirate or nasal aspirate specimens, and BAL specimens | Applied Biosystems QuantStudio 12K Flex Real-Time PCR System (QuantStudio 12K Software v1.3) | Positive for two or more of the target genes |  | https://www.fda.gov/media/142535/download
09/29/2020 | CENTOGENE US, LLC | CentoSure SARS-CoV-2 RT-PCR Assay | No | H | N1, N2 | 2 viral copies/µL | N/A | 2-4 hours | Dry OP swab specimens | LightCycler 480 II is using Software Version 1.5.1.62 or higher | Positive for both of the target genes |  | https://www.fda.gov/media/142539/download
09/29/2020 | Akron Children’s Hospital | Akron Children’s Hospital SARS-CoV-2 Assay | No | H | S, E | 1250 copies/mL | N/A | 2-4 hours | NPl swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and BAL fluid specimens | QIAGEN RotorGene Q MDx PCR instrument (Software V2.1.0.9) | Positive for S with or without E | E gene is specific for lineage B-β-coronaviruses | https://www.fda.gov/media/142543/download
09/25/2020 | Genetrack Biolabs, Inc. | Genetrack SARS-CoV-2 Molecular Assay | No | H | N1, N2 | 1 GC/µL | N/A | 2-4 hours | Anterior nasal swab specimens | Applied Biosystems QuantStudio 6 Pro Real-Time PCR System equipped with software v2.4.1 | Positive for both of the target genes |  | https://www.fda.gov/media/142482/download
09/22/2020 | Quadrant Biosciences Inc. | Clarifi COVID-19 Test Kit | Yes | H | 2 regions of RdRp | 1 copy/µL | N/A | 2-4 hours | Saliva swab specimens | BioRad-CFX96 Touch Real Time Detection System, BioRad-CFX384 Touch Real Time Detection System, QuantStudio 5 | Positive for SARS-CoV-2 HEX | The detection of SARS-CoV-2 RNA is performed using two target regions, which are both detected in the same fluorescence channel (HEX) | https://www.fda.gov/media/142379/download
09/22/2020 | Vela Operations Singapore Pte. Ltd. | ViroKey SARS-CoV-2 RT-PCR Test v2.0 | Yes | H | Orf1a, N | 200 GE/mL | N/A | 2-4 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal or NP aspirates, nasal washes and BAL samples | Applied Biosystems® 7500 Fast Dx Real-Time PCR System or the Sentosa® SA201 Real-Time PCR Instrument | Positive for one or both of the target genes |  | https://www.fda.gov/media/142383/download
09/21/2020 | KimForest Enterprise Co., Ltd. | KimForest SARS-CoV-2 Detection Kit v1 | Yes | H | RdRp | 200 copies/mL | N/A | 2-4 hours | NP/OP swabs, anterior/mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates, and BAL specimens | NR | Positive for RdRp |  | https://www.fda.gov/media/142363/download
09/18/2020 | Infinity BiologiX LLC | Infinity BiologiX TaqPath SARS-CoV-2 Assay | No | H | Orf1ab, S, N | 200 copies/mL | N/A | N/A | OP swab, NP swab, anterior nasal swab, mid-turbinate nasal swab, and BAL fluid | ThermoFisher Applied Biosystems QuantStudio 5 Real-Time PCR System equipped with software v1.3, or the Applied Biosystems ViiA7 Real-Time PCR System with the Applied Biosystems QuantStudio 5 software v1.3 | Positive for two or more of the target genes |  | https://www.fda.gov/media/137776/download
09/18/2020 | GK Pharmaceuticals Contract Manufacturing Operations | GK ACCU-RIGHT SARS-CoV-2 RT-PCR KIT | Yes | H | N1, N2 | 2 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (such as NP, OP, mid-turbinate, and nasal swabs) | Applied Biosystems™ 7500 Fast Dx Real-Time PCR Instrument with SDS software version 1.4 | Positive for one or both of the target genes |  | https://www.fda.gov/media/142310/download
2020-09-09 | Beijing Wantai Biological Pharmacy Enterprise Co., Ltd. | Wantai SARS-CoV-2 RT-PCR Kit | Yes | H | Orf1ab, N | 50 copies/mL | N/A | 2-4 hours | OP swab specimen | BIO-RAD CFX-96 (Bio-Rad CFX Manager 3.1), ABI7500 (software v2.3) | Positive for both of the target genes |  | https://www.fda.gov/media/141997/download
2020-08-09 | Verily Life Sciences | Verily COVID-19 RT-PCR Test | No | H | Orf1ab, S, N | 60 GCE/mL | N/A | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) | 7500 Dx Fast Real-Time PCR with SDS Software v1.4.1 or QuantStudio 5 Real-Time PCR Instrument 384-well block with QuantStudio Design and Analysis Desktop Software v1.5.1 | Positive for two or more of the target genes |  | https://www.fda.gov/media/141951/download
2020-02-09 | Bioeksen R&D Technologies Ltd. | Bio-Speedy Direct RT-qPCR SARS-CoV-2 | Yes | H | Orf1ab | 200 copies/mL (NP aspirate); 281 copies/mL (BAL); 562 copies/mL (OP); 89 copies/mL (NP) | N/A | 2-4 hours | NP swabs, OP swabs, combined NP/OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal or NP aspirates, nasal washes and BAL samples | Roche LightCycler® 96, Bio-Rad CFX96 Touch™, Qiagen Rotor-Gene® 5 Plex Real-Time PCR Systems | Positive for Orf1ab |  | https://www.fda.gov/media/141823/download
2020-01-09 | OPTOLANE Technologies, Inc. | Kaira 2019-nCoV Detection Kit | Yes | H | RdRp, E | 2500 copies/mL (E); 5000 copies/mL (RdRp) | N/A | 2-4 hours | NP, OP, anterior nasal, and mid-turbinate nasal swab specimens, as well as NP wash/aspirate, nasal aspirate, sputum and BAL specimens | QuantStudio 5 Real-Time PCR system (Thermo-Fisher Scientific; Software ver. 1.5.1), ABI 7500 Real-Time PCR system (Applied Biosystems co.; Software ver. 2.3), or CFX96 Dx System (Bio-Rad Laboratories; Software ver. 3.1.1517.0823) | Positive for RdRp with or without E |  | https://www.fda.gov/media/141795/download
08/31/2020 | MiraDx | MiraDx SARS-CoV-2 RT-PCR assay | No | H | N1, N2 | 4 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, OP swab specimens and NP wash/aspirate or nasal aspirate specimens) | Applied Biosystems QuantStudio 6 Flex PCR Instrument | Positive for one or both of the target genes |  | https://www.fda.gov/media/141760/download
08/28/2020 | Patients Choice Laboratories, LLC | PCL SARS-CoV-2 Real-Time RT-PCR Assay | No | H | Orf1ab, S, N | 100 copies/µL | 5400 | N/A | NP and OP swabs | Applied Biosystems QuantStudio 12K Flex ( QuantStudio 12K Flex Gene Expression software v1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/141665/download
08/26/2020 | Cuur Diagnostics | Cuur Diagnostics SARS-CoV-2 Molecular Assay | No | H | Orf1ab, S, N | 25 copies/µL | N/A | N/A | Upper respiratory specimens (such as NP, OP, mid-turbinate, and nasal swab) | Applied Biosystems QuantStudio 12K Flex ( QuantStudio 12K Flex Gene Expression software v1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/141627/download
08/25/2020 | Fluidigm Corporation | Advanta Dx SARS-CoV-2 RT-PCR Assay | Yes | H | N1, N2 | 6.25 GE/µL | 54000 | 2-4 hours | Saliva specimens | Biomark HD system with Data Collection software v4.5.2 | Positive for one or both of the target genes |  | https://www.fda.gov/media/141541/download
08/21/2020 | DxTerity Diagnostics, Inc. | DxTerity SARS-CoV-2 RT-PCR Test | No | H | Orf1ab, S, N | 50 copies/mL | N/A | 2-4 hours | Saliva specimens | Thermo Fisher Applied Biosystems ViiA 7 Real-Time PCR amplification system (QuantStudio5 software v1.3) | Positive for two or more of the target genes |  | https://www.fda.gov/media/141491/download
08/21/2020 | Texas Department of State Health Services, Laboratory Services Section | Texas Department of State Health Services (DSHS) SARS-CoV-2 Assay | No | H | Orf1ab, N | 20 copies/mL | N/A | 2-4 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and BAL fluid specimens | PerkinElmer Chemagic 360 (software version 6.3.0.3), ThermoFisher Applied Biosystems 7500 Fast Dx Real-Time PCR System (software version 1.4) | Positive for one or both of the target genes |  | https://www.fda.gov/media/141496/download
08/17/2020 | ZhuHai Sinochips Bioscience Co., Ltd. | COVID-19 Nucleic Acid RT-PCR Test Kit | Yes | H | Orf1ab, N | 2 copies/µL | 5400 | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, OP swab specimens and NP wash/aspirate or nasal aspirate specimens) | Applied BiosystemsTM ABI 7500 Real-Time PCR system (Software v2.3) or Applied BiosystemsTM 7500 Fast Dx Real-Time PCR Instrument (Software v1.4.1) | Positive for both of the target genes |  | https://www.fda.gov/media/141246/download
08/15/2020 | Yale School of Public Health, Department of Epidemiology of Microbial Diseases | SalivaDirect | No | H | N1 | 6 copies/µL | 18000 | N/A | Saliva specimens | CFX96 Touch Real-Time PCR Detection System, Applied Biosystems 7500 Fast Real-Time PCR System, Applied Biosystems 7500 Fast Dx Real-Time PCR System | Positive for N1 |  | https://www.fda.gov/media/141192/download
2020-11-08 | Biomeme, Inc. | Biomeme SARS-CoV-2 Real-Time RT-PCR Test | Yes | H | Orf1ab, S | 1.8 GE/µL | N/A | 2-4 hours | NP, nasal, and OP swab specimens, and NP wash/aspirate or nasal aspirate specimens | Biomeme’s portable Franklin™ Real-Time PCR Thermocycler, Bio-Rad CFX96 or QuantStudio5 | Positive for S with or without Orf1ab |  | https://www.fda.gov/media/141052/download
2020-10-08 | Solaris Diagnostics | Solaris Multiplex SARS-CoV-2 Assay | No | H | N1, N2 | 10 copies/µL | N/A | 2-4 hours | Respiratory specimens (NP swabs, anterior nasal swabs, mid-turbinate nasal swabs, OP swabs, NP wash/aspirates, nasal aspirates, BAL specimens) | Applied Biosystems QuantStudio 5 (Software version v1.5.1) | Positive for both of the target genes |  | https://www.fda.gov/media/141016/download
2020-10-08 | Alpha Genomix Laboratories | Alpha Genomix TaqPath SARS-CoV-2 Combo Assay | No | H | Of1ab, S, N | 4 copies/µL | N/A | 2-4 hours | NP, OP, anterior nasal, and mid-turbinate nasal swab specimens | QuantStudio 12K Flex Real-Time PCR System (Software version 2.4.0) | Positive for two or more of the target genes |  | https://www.fda.gov/media/141021/download
2020-07-08 | George Washington University Public Health Laboratory | GWU SARS-CoV-2 RT-PCR Test | No | H | N1, N2 | 12.5 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP and OP swabs) and BAL | LightCycler 480 instrument (Software version 1.5) | Positive for both of the target genes |  | https://www.fda.gov/media/140980/download
2020-05-08 | Vela Operations Singapore Pte Ltd | ViroKey SARS-CoV-2 RT-PCR Test | Yes | H | Orf1a, RdRp | 560 GE/mL | N/A | 2-4hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, OP swab specimens and nasopharyngeal wash/aspirate or nasal aspirate specimens) | Applied Biosystems® 7500 Fast Dx Real-Time PCR System or the Sentosa® SA201 Real-Time PCR Instrument | Positive for one or both of the target genes |  | https://www.fda.gov/media/140922/download
2020-03-08 | Wren Laboratories LLC | Wren Laboratories COVID-19 PCR Test | No | H | N1, N3 | 10 copies/µL | N/A | 2-4 hours | NP, OP, anterior nasal, and mid-turbinate nasal swabs, as well as NP washes/aspirates or nasal aspirates, and BAL specimens | Applied Biosystems QuantStudio 7-Flex Real-Time PCR Instrument (QuantStudio Real-Time PCR software v1.3) | Positive for N1 with or without N3 | N3 target is specific to Sarbecovirus/SARS-like coronaviruses | https://www.fda.gov/media/140776/download
2020-03-08 | Cleveland Clinic Robert J. Tomsich Pathology and Laboratory Medicine Institute | Cleveland Clinic SARS-CoV-2 Assay | No | H | E, RdRp | 10 copies/µL | N/A | 2-4 hours | NP, OP and nasal swabs and BAL, tracheal aspirate (TASP), and sputum | Roche LightCycler 480 and/or cobas z 480 instruments (software version v1.2.1.62 SP3) | Positive for both of the target genes |  | https://www.fda.gov/media/140788/download
2020-03-08 | ISPM Labs, LLC dba Capstone Healthcare | Genus SARS-CoV-2 Assay | No | H | N1, N2 | 40 copies/µL | N/A | N/A | NP and OP swabs | Applied Biosystems QuantStudio 12K Flex Real Time PCR Instrument (QuantStudio RealTime PCR System software version 1.3) | Positive for both of the target genes |  | https://www.fda.gov/media/140818/download
07/31/2020 | Xiamen Zeesan Biotech Co., Ltd. | SARS-CoV-2 Test Kit (Real-time PCR) | Yes | H | Orf1ab, N | 200 copies/mL | N/A | 2-4 hours | NP/OP swabs, anterior/mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates, and BAL specimens | Applied Biosystems™ QuantStudio 3 real-time PCR thermal cycler (QuantStudioTM Design & Analysis Software v1.4.3) | Positive for both of the target genes |  | https://www.fda.gov/media/140717/download
07/30/2020 | Clinical Reference Laboratory, Inc | CRL Rapid Response | No | H | RdRp | 0.25 viral copies/µL | 5400 | N/A | Saliva specimens | BioRad CFX-96 TOUCH Thermal Cycler instrument (CFX Manager 3.1 software) | Positive for RdRp |  | https://www.fda.gov/media/140661/download
07/27/2020 | Eli Lilly and Company | Lilly SARS-CoV-2 Assay | Yes | H | N1, N1 | 103 copies of packaged viral RNA/mL | 18000 | 2-4 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and BAL fluid | Applied Biosystems QuantStudio 7 Flex Real-Time PCR System (software v1.3), Applied Biosystems QuantStudio Dx Real-Time PCR System (software v1.0.3), Applied Biosystems QuantStudio 12K Flex RealTime PCR System (software v1.3) | Positive for one or both of the target genes |  | https://www.fda.gov/media/140543/download
07/27/2020 | Sandia National Laboratories | SNL-NM 2019 nCoV Real-Time RT-PCR Diagnostic Assay | No | H | N1, N2 | 6.25 copies/µL | N/A | 2 hours | Upper respiratory specimens (including NP swab, anterior nasal swab, mid-turbinate nasal swab and OP swab, nasal washes, nasal aspirates) and BAL specimens | Applied Biosystems 7500 (ABI7500) (7500 Software v2.3), Applied Biosystems Quant Studio 5 (Software v1.5.1), Bio-Rad CFX Connect (CFX Manager v3.1), or Bio-Rad CFX 96 Real PCR instruments (CFX Maestro 1.1 V4.1.2433.1219) | Positive for both of the target genes |  | https://www.fda.gov/media/140547/download
07/24/2020 | Jiangsu CoWin Biotech Co., Ltd. | Novel Coronavirus (SARS-CoV-2) Fast Nucleic Acid Detection Kit (PCR-Fluorescence Probing) | Yes | H | Orf1ab, N | 300 copies/mL | N/A | 2 hours | OP swab | Applied Biosystems™ 7500 Real-time PCR System | Positive for both of the target genes |  | https://www.fda.gov/media/140425/download
07/23/2020 | Helix OpCo LLC (dba Helix) | Helix COVID-19 Test | No | H | Orf1ab, S, N | 1000 GCE/mL | 1800 | 2 hours | Upper respiratory specimens (NP swabs, OP swab, mid-turbinate nasal swabs and anterior nasal swabs) | QuantStudio 7 Flex Quantitative Real Time PCR Instrument (Software v2.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/140420/download
07/21/2020 | DiaCarta, Inc. | QuantiVirus SARS-CoV-2 Multiplex Test Kit | Yes | H | Orf1ab | 100 copies/mL (QS5, CFX384), 50 copies/mL (ABI 7500 Fast Dx) | 600 | 2 hours | Upper respiratory specimens (such as nasal, NP and OP swab specimens) and sputum | Applied Biosystems™ QuantStudio 5 Real-Time PCR Instrument (QuantStudio™ Design and Analysis Software v1.4), ABI 7500 Fast DX (SDS Software v1.4) and Bio-Rad CXF 384 Real-Time PCR Instrument (Bio-Rad CFX Manager v3.1) | Positive for Orf1ab |  | https://www.fda.gov/media/140354/download
07/17/2020 | Access Genetics, LLC | OraRisk COVID-19 RT-PCR | No | H | RdRp | 15 copies/µL | N/A | 2 hours | NP swab and nasal swab specimens | LightCycler 480 II instrument (Software version 1.5.1.62) | Positive for RdRp |  | https://www.fda.gov/media/140293/download
07/16/2020 | Boston Heart Diagnostics | Boston Heart COVID-19 RT-PCR Test | No | H | Orf1ab, S, N | 250 copies/mL | 5400 | 2 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP and OP swab specimens), NP aspirate and BAL specimens | QuantStudio 12K Flex Real-Time PCR System (Software version 1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/140259/download
07/13/2020 | KogeneBiotech Co., Ltd. | PowerChek 2019-nCoV Real-time PCR Kit | Yes | H | E, RdRP | 4 copies/µL | 5400 | 2-4 hours | Anterior/mid-turbinate nasal swabs, NP/OP swabs, nasopharyngeal washes/aspirates or nasal aspirates, BAL, and sputum | Applied Biosystems™ 7500 Real-time PCR System, Applied Biosystems™ 7500 fast System run as a standard AB 7500, CFX96™ Real-time PCR Detection System | Positive for RdRP with or without E | E gene is specific for Sarbecovirus | https://www.fda.gov/media/140069/download
07/13/2020 | Trax Management Services Inc. | PhoenixDx SARS-CoV-2 Multiplex | Yes | H | Orf1ab, N | 50 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP and OP swabs) and BAL specimens | Thermo Fisher QS1, Qiagen Rotor-Gene Q, Applied Biosystems ABI 7500 Fast Real time PCR Dx, DNA Technologie DTPrime5, and Analytik Jena qTower3G platforms | Positive for SARS-CoV-2 FAM | The detection of SARS-CoV-2 RNA is performed using two target regions, which are both detected in the same fluorescence channel (FAM) | https://www.fda.gov/media/140073/download
07/13/2020 | Compass Laboratory Services, LLC | Compass Laboratory Services SARS-CoV2 Assay | No | H | Orf1ab | 5 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (NP swabs, OP swab, mid-turbinate nasal swabs, anterior nasal swabs, NP wash/aspirate or nasal aspirate) | QuantStudio 12K Flex (software v1.3) and QuantStudio 7 Flex Quantitative Real Time PCR Instruments (Software v1.3) | Positive for Orf1ab |  | https://www.fda.gov/media/140078/download
2020-10-07 | Boston Medical Center | BMC-CReM COVID-19 Test | No | H | N1, N2 | 10 gene copies/µL (Two-Step Singleplex Test), 1 gene copy/µL (One-Step Singleplex Test) | 600000 | 2 hours | NP, OP, anterior nasal and mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates, and BAL specimens | QuantStudio 6 Flex Real-Time PCR System (software version 1.3) | Positive for one or both of the target genes |  | https://www.fda.gov/media/140007/download
2020-09-07 | BioSewoom, Inc. | Real-Q 2019-nCoV Detection Kit | Yes | H | E, RdRP | 6.25 copies/µL (sputum), 6.25 copies/µL (NP swab using ABI7500, QS5, CFX96), 3.125 copies/µL (NP swab using ABI7500 fast and CFX96) | N/A | N/A | NP swab, NP aspirates, nasal mid-turbinate swab, nasal swab, OP swab, sputum, tracheal aspirates and BAL specimens | Applied Biosystems 7500 real-time PCR instrument (7500 Software v2.3), Applied Biosystems 7500 fast real-time PCR instrument (7500 Software v2.3), Applied Biosystems QuantStudio 5 real-time PCR Instrument (QuantStudio™ Design and Analysis Software v1.4), CFX96 real-time PCR detection system (CFX Manager™ Software v1.6), CFX96 DX real-time PCR detection system (CFX Manager™ Software v1.6) | Positive for RdRP with or without E |  | https://www.fda.gov/media/139941/download
2020-08-07 | Clinical Research Sequencing Platform (CRSP), LLC at the Broad Institute of MIT and Harvard | CRSP SARS-CoV-2 Real-time Reverse Transcriptase (RT)-PCR Diagnostic Assay | No | H | Assay version 1: N1, N2 Assay version 2: N2 | Assay version 1: 4.0 x 103 copies/mL Assay version 2: 1.6 x 103 copies/mL | N/A | ~2 Hours | Upper respiratory specimens (such as NP, OP, nasal, and mid-turbinate swabs, NP wash/aspirate or nasal aspirate specimens) and BAL specimens | Assay version 1: Applied Biosystems Viia7 thermocycler or the Applied Biosystems QuantStudio 7 Flex each with QuantStudio software version 1.3 Assay version 2: Applied Biosystems QuantStudio 7 Flex thermocycler with QuantStudio software version 1.3 | Assay version 1: Positive for both N1 and N2 Assay version 2: Positive for N2 |  | https://www.fda.gov/media/139858/download
2020-07-07 | Enzo Life Sciences, Inc. | AMPIPROBE SARS-CoV-2 Test System | Yes | H | N1, N2 | 280 copies/mL | 18000 | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, OP swab specimens and NP wash/aspirate or nasal aspirate specimens) | Applied Biosystems QuantStudio® 5 (QS5) instrument (Software version 1.5.1) | Positive for both N1 and N2 |  | https://www.fda.gov/media/139828/download
2020-07-07 | Access Bio, Inc. | CareStartTM COVID-19 MDx RT-PCR | Yes | H | N, RdRP | 10 copies/reaction | 5400 | 2-4 hours | Upper respiratory specimens (such as NP, OP and nasal swabs, and NP wash/aspirate or nasal aspirate) and BAL | Applied Biosystems™ 7500/7500 Fast Real-Time PCR Instrument (SDS Software v1.4. 21 CFR Part 11 Module), CFX96 Touch™ Real-Time PCR Detection System (CFX Manager™ Software v3.1) | Positive for one or both of the target genes |  | https://www.fda.gov/media/139832/download
2020-07-07 | Gene By Gene | Gene By Gene SARS-CoV-2 Detection Test | No | H | N1, N2 | 6.25 copies/µL | N/A | 2-4 hours | Nasal swabs and NP swabs | Bio-Rad CFX384 Touch Real-Time PCR Detection System (Bio-Rad CFX Maestro Software ) | Positive for both N1 and N2 |  | https://www.fda.gov/media/139836/download
2020-06-07 | Laboratorio Clinico Toledo | Laboratorio Clinico Toledo SARS-CoV-2 Assay | No | H | E, N, RdRP | 203 copies/mL (RdRp and E); 21000 copies/mL (N) | N/A | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) and BAL specimens | Roche Light Cycler 480 (software version LCS480 1.5.16.2 SP@-UDF v 2.0) | Positive for RdRp, N and E or RdRP with/without N or E |  | https://www.fda.gov/media/139788/download
2020-02-07 | Centers for Disease Control and Prevention (CDC) | Influenza SARS-CoV-2 (Flu SC2) Multiplex Assay | Yes | H | N | 1.01×10-2 TCID50/mL (TaqPath 1-step Multiplex); 5.06×10-2 TCID50/mL (Ultraplex 1-step ToughMix) | 5700 | 2-4 hours | Upper or lower respiratory specimens (such as NP, OP and nasal swabs, sputum, lower respiratory tract aspirates, BAL, and NP wash/aspirate or nasal aspirate) | Applied Biosystems 7500 Fast Dx Real-Time PCR Instrument (SDS version 1.4.1 software) | Positive for N |  | https://www.fda.gov/media/139743/download
2020-01-07 | Centogene US, LLC | CentoFast-SARS-CoV-2 RT-PCR Assay | No | H | E, RdRP | 5 viral copies/µL | N/A | 2-4 hours | Dry OP swab | QuantStudio 5 Real-Time PCR System, LightCycler 480 II (Roche) (Software Version 1.5.1.62 or higher) | 1st Tier: Positive for E 2nd Tier: Positive for both E and RdRP |  | https://www.fda.gov/media/139725/download
06/30/2020 | Psomagen, Inc. | Psoma COVID-19 RT Test | No | H | N1, N2 | 1 copy/µL | 18000 | 2 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) and BAL specimens | Light cycler 480 instrument (Light Cycler 480 software Version 1.5) | Positive for both N1 and N2 |  | https://www.fda.gov/media/139676/download
06/30/2020 | TNS Co., Ltd (Bio TNS) | COVID-19 RT-PCR Peptide Nucleic Acid (PNA) kit | Yes | H | RdRP, N | 3.715 copies/µL (RdRp); 2.524 copies/µL (N) | N/A | N/A | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, BAL, and NP wash/aspirates or nasal aspirates | Bio-Rad CFX-96 real-time PCR detection system | Positive for one or both of the target genes |  | https://www.fda.gov/media/139680/download
06/29/2020 | LifeHope Labs | LifeHope 2019-nCoV Real-Time RT-PCR Diagnostic Panel | No | H | N1, N2 | 2.5 GE/µL | N/A | ~2 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) and BAL specimens | QuantStudio 5 (Software version 1.5.1) | Positive for both N1 and N2 |  | https://www.fda.gov/media/139623/download
06/29/2020 | Acupath Laboratories, Inc. | Acupath COVID-19 Real-Time (RT-PCR) Assay | No | H | Orf1ab, S, N | 25 copies/µL | 18000 | 2 hours | NP swab, NP aspirate, and BAL specimens | Applied Biosystems QuantStudio 12K Flex (QuantStudio software version 1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/139672/download
06/26/2020 | Inform Diagnostics, Inc. | Inform Diagnostics SARS-CoV-2 RT-PCR Assay | No | H | N1, N2 | 20 copies/µL | N/A | 1 hour | NP, OP, anterior nasal, and midturbinate nasal swabs, as well as NP wash/aspirate or nasal aspirates, and BAL specimens | 7500 Fast Dx Real-Time PCR Instrument ( software version 1.4) or QuantStudio (software version 1.0.3) | Positive for either N1 or N2 or both |  | https://www.fda.gov/media/139572/download
06/25/2020 | Diagnostic Solutions Laboratory, LLC | DSL COVID-19 Assay | No | H | S, N1, N3 | 10 copies/swab | N/A | 2 hours | Upper respiratory specimens (such as nasal, midturbinate, NP, and OP swabs) and BAL specimens | QuantStudio 5 Real-Time PCR System (QS5 Analysis software) | Postitive for N1 and/or S | N3 is a coronavirus generic primer and probe set that can also detect other coronaviruses | https://www.fda.gov/media/139516/download
06/24/2020 | University of Texas MD Anderson Cancer Center, Molecular Diagnostics Laboratory | MD Anderson High-throughput SARS-CoV-2 RT-PCR Assay | No | H | N1, N2 | 5 copies/µL | N/A | 1.5 -2 hours | Upper respiratory specimens (such as NP, OP, nasal, and mid-turbinate swabs) | QuantStudio 7 Flex Fast Sequence Detection System (Software V1.3), QuantStudio Dx Real-Time PCR Instrument (Software V1.0.3), Applied Biosystems 7500 Fast Real-Time PCR Instrument (Software V1.4) | Positive for both N1 and N2 |  | https://www.fda.gov/media/139512/download
06/23/2020 | HealthQuest Esoterics | HealthQuest Esoterics TaqPath SARS-CoV-2 Assay | No | H | Orf1ab, S, N | 20 copies/µL | 18000 | <3 hours | NP, OP, anterior nasal, and mid-turbinate nasal swabs, as well as NP wash/aspirate or nasal aspirates, and BAL specimens | QuantStudio 12K Flex Real-Time PCR System (software version 1.3) | Positive for one or more of the target genes |  | https://www.fda.gov/media/139434/download
06/23/2020 | University of Alabama at Birmingham Fungal Reference Lab | FRL SARS CoV-2 Test | No | H | N1 | 125 copies/mL | N/A | <3 hours | NP swab, OP swab, nasal swab, mid-turbinate nasal swab, anterior nasal swab,NP wash/aspirate, nasal aspirate, and BAL specimens | Applied Biosystems QuantStudio 5 Real-Time PCR System (Software v1.4.3) | Positive for N1 |  | https://www.fda.gov/media/139437/download
06/23/2020 | Gencurix, Inc | GenePro SARS-CoV-2 Test | Yes | H | Orf1ab, E | 5550 GE/mL (Orf1ab) 6,000 GE/mL (E ) | N/A | 2 hours | NP swab, OP swab, anterior nasal swab, and mid-turbinate swab specimens as well as NP wash/ aspirate, nasal aspirate and BAL specimens | Applied Biosystems™ QuantStudio™ Dx (Software v1.0.3), CFX96™ Dx System (Software v3.1) | Positive for Orf1ab with or without E |  | https://www.fda.gov/media/139442/download
06/18/2020 | Jiangsu Bioperfectus Technologies Co., Ltd. | COVID-19 Coronavirus Real Time PCR Kit | Yes | H | Orf1ab, N | 3.5×102 copies/mL | N/A | ~2 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes, BAL fluid and sputum | Applied Biosystems 7500 (Software V2.3 and V2.4), QuantStudio™ 5 (Software version V1.4.3 and V1.5.1), Roche LightCycler® 480 (Software version V1.5.1.62) | Positive for one or both of the target genes |  | https://www.fda.gov/media/139279/download
06/18/2020 | 3B Blackbio Biotech India Ltd., a subsidiary of Kilpest India Ltd. | TRUPCR SARS-CoV-2 Kit | Yes | H | E, N, RdRp | 10 copies/µL | N/A | 1.5-2 hours | NP and OP swabs, anterior nasal swab and mid-turbinate nasal swabs, NP aspirates/washes or nasal aspirates, and BAL specimens | Applied Biosystem QuantStudio 3 (Software version 1.4.3), Qiagen Rotor-Gene Q 5Plex HRM Platform (software version 2.3.1) | Positive for SARS-CoV-2 FAM | E gene for the detection of the Sarbecovirus (of Genus B-betacoronavirus (B-βCoV)) | https://www.fda.gov/media/139296/download
06/17/2020 | The Ohio State University Wexner Medical Center | OSUWMC COVID-19 RT-PCR test | No | H | N1, N2 | 0.25 copies/µL | N/A | 30 minutes | Nasal, OP and NP swabs | ABI 7500 Fast Real-time PCR instrument (software version 1.5.1) | Positive for one or both of the target genes |  | https://www.fda.gov/media/139288/download
2020-12-06 | RTA Laboratories Biological Products Pharmaceutical and Machinery Industry | Diagnovital SARS-CoV-2 Real-Time PCR Kit | Yes | H | E, RdRp | 38 copies/mL | N/A | N/A | Anterior nasal and mid-turbinate nasal swabs, NP and OP swabs, NP wash/aspirates or nasal aspirates and BAL specimens | BIO-RAD CFX96-IVD (Bio-Rad Laboratories, Inc.) Software version 3.0 and QuantStudio™ 5 Dx, (Applied Biosystems) Software v 1.0 | Positive for RdRp with or without E |  | https://www.fda.gov/media/138928/download
2020-10-06 | Tide Laboratories, LLC | DTPM COVID-19 RT-PCR Test | Yes | H | N | 22 copies/µL | N/A | 25 minutes | NP swabs, OP swabs, nasal swabs, and mid-turbinate swab specimens | ThermoFisher QuantStudio 5 (software version 1.5.1) | Positive for N |  | https://www.fda.gov/media/138818/download
2020-10-06 | TBG Biotechnology Corp. | ExProbe SARS-CoV-2 Testing Kit | Yes | H | E, N, RdRp | 10 copies/µL | N/A | ~35 minutes | NP and OP swabs, anteri or nasal and mid-turbinate nasal swabs, NP aspirate/wash and nasal aspirate specimens, and BAL specimens | ABI 7500 Real Time PCR System (Software version 2.3) | Positive for RdRp with or without E and N | E and N are for the detection of Sarbecovirus | https://www.fda.gov/media/138822/download
2020-08-06 | Euroimmun US, Inc. | EURORealTime SARS-Cov-2 | Yes | H | Orf1ab, N | 150 copies/mL | 1800 | 1.5 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swabs) and bronchioalveolar lavage (BAL) | LightCycler® 480 II, 7500 Fast Real-Time PCR Instrument, CFX 96 Touch | Positive for SARS-CoV-2 FAM | The detection of SARS-CoV-2 RNA is performed using two target regions, which are both detected in the same fluorescence channel (FAM) | https://www.fda.gov/media/138761/download
2020-05-06 | Genetron Health (Beijing) Co., Ltd. | Genetron SARS-CoV-2 RNA Test | Yes | H | Orf1ab, N | 1000 copies/mL | N/A | ~2.5 hours | Upper respiratory specimens (such as OP, NP, anterior nasal and mid-turbinate nasal swab specimens) | Applied Biosystems 7500 Real-Time PCR System (software version 2.3) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138685/download
2020-04-06 | Phosphorus Diagnostics LLC | Phosphorus COVID-19 RT-qPCR Test | No | H | N1, N2 | 5 copies/µL | 18000 | Unavailable | OP swabs, NP swabs, anterior nasal and mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates as well as BAL specimens | CFX384 Touch Real-Time PCR Detection System ( Software version 3.1) | Positive for both N1 and N2 |  | https://www.fda.gov/media/138654/download
2020-01-06 | Gravity Diagnostics, LLC | Gravity Diagnostics COVID-19 Assay | No | H | N1, N2 | 2.4 copies/µL | 18000 | 2 hours | Nasal, NP and OP swab and BAL specimens | Applied Biosystems QuantStudio7 Flex instrument (software version 1.3) and QuantStudio12 Flex instrument (software version 1.2.2) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138530/download
05/22/2020 | dba SpectronRX | Hymon SARS-CoV-2 Test Kit | Yes | H | E, N | 5 copies/µL | N/A | N/A | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) and BAL specimens | ABI 7500 Real-Time PCR System | Positive for SARS-CoV-2 FAM |  | https://www.fda.gov/media/138345/download
05/22/2020 | Exact Sciences Laboratories | SARS-CoV-2 (N gene detection) Test | No | H | N1 | 1.2 copies/µL | 6025 | N/A | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) | ABI 7500 Fast Dx (Software version 1.4.1) | Positive for N1 |  | https://www.fda.gov/media/138328/download
05/21/2020 | P23 Labs, LLC. | P23 Labs TaqPath SARS-CoV-2 Assay | No | H | Orf1ab, S, N | 10 copies/µL | N/A | Unavailable | OP swabs, NP swabs, anterior nasal and mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates as well as BAL specimens | ThermoFisher Applied Biosystems QuantStudio 5 Real-Time PCR System (software v1.3 ) | Positive for one or more of the target genes |  | https://www.fda.gov/media/138297/download
05/21/2020 | SolGent Co., Ltd | DiaPlexQ Novel Coronavirus (2019-nCoV) Detection Kit | Yes | H | Orf1a, N | 200 copies/mL | N/A | ~2 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes, BAL fluid and sputum | Applied Biosystems™ 7500 Real-Time PCR Instrument System and 7500 Fast Real-Time PCR Instrument System (software V2.0.6), Bio-Rad CFX96™ Real-time PCR Detection System (software V3.1) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138303/download
05/21/2020 | BioCore Co., Ltd. | BioCore 2019-nCoV Real Time PCR Kit | Yes | H | N, RdRP | 500 copies/ mL | 600 | 2 hours | Upper respiratory specimens (such as nasal, mid-turbinate, OP, and NP swabs) and lower respiratory specimens (such as sputum, bronchioalveolar lavage (BAL), and tracheal aspirates) | CFX96 Dx System (Bio-Rad Inc.), Applied Biosystems 7500 RealTime PCR Instrument System (Thermo Fisher Scientific Inc.), or SLAN-96P (Shanghai Hongshi Medical Technology Co. Ltd) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138290/download
05/18/2020 | Quidel Corporation | Lyra Direct SARS-CoV-2 Assay | Yes | H | pp1ab | 1.28×104 GE/mL | 540000 | 2.4 hours | Nasal, NP, or OP direct swab specimens | Applied Biosystems 7500 Fast Dx (software version 1.4, or later), Applied Biosystems 7500 Standardsoftware version 2.0.6, or later), Roche LightCycler 480 (software version 1.5.0.39, or later), Roche Cobas z480 (software version 1.5.1.62 SP2-, or later), Qiagen Rotor-Gene Q (software version 2.0.2.4, or later), Bio-Rad CFX96 Touch (software version 3.1, or later), Thermofisher QuantStudio 7 Pro (software version 2.4, or later) | Positive for pp1ab |  | https://www.fda.gov/media/138178/download
05/15/2020 | Fulgent Therapeutics, LLC | Fulgent COVID-19 by RT-PCR Test | No | H | N1, N2 | 5 copies/µL | 3600 | 1-2 days | Upper and lower respiratory specimens (Nasal swabs, NP and OP swabs) | QuantStudio 6 and QuantStudio 7 Real-Time PCR System | Positive for both N1 and N2 |  | https://www.fda.gov/media/138150/download
05/15/2020 | Assurance Scientific Laboratories | Assurance SARS-CoV-2 Panel | No | H | N1 | 5 GC/µL (Abnova); 9 GC/µL (Zymo) | 5400 | N/A | Nasal swabs, NP and OP swabs | BioRad CFX96 or CFX384 | Positive for both N1 and N2 |  | https://www.fda.gov/media/138154/download
05/14/2020 | GeneMatrix, Inc. | NeoPlex COVID-19 Detection Kit | Yes | H | N, RdRp | 50 copies/reaction | 5400 | 3-4 hours | Upper respiratory specimens (such as NP, OP, mid-turbinate, and nasal swabs) and lower respiratory specimens (such as sputum, tracheal aspirates, and BAL) | Applied Biosystems™ 7500 Fast Dx Real-Time PCR Instrument (software version 1.4) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138100/download
05/13/2020 | Applied DNA Sciences, Inc. | Linea COVID-19 Assay Kit | Yes | H | S1, S2 | 1.25 copies/µL (TRIzol and QIAamp extraction kits), 0.625 copies/µL (Omega Bio-Tek Mag-Bind Viral RNA Xpress Kit) | 2500 | 1 hour | upper respiratory specimens (such as NP swabs and OP swabs, anterior nasal and mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates, and BALs | Thermo Fisher Scientific (Applied Biosystems) QuantStudioTM Dx Real-Time PCR system (software v1.0.3) and QuantStudioTM 5 Real-Time PCR system (software v1.4) | Positive for one or both of the target genes |  | https://www.fda.gov/media/138059/download
2020-11-05 | 1drop Inc. | 1copy COVID-19 qPCR Multi Kit | Yes | H | E, RdRp | 200 copies/mL | 18000 | <2 hours | NP swab, anterior nasal swab, mid-turbinate nasal swab and OP swab as well as NP wash/aspirate and nasal aspirate specimens | Light Cycler 480 (Software version 1.5), Rotor-Gene Q 5plex HRM (Software version 2.3.4), Applied Biosystems Quantstudio5 (Software version 1.4.3), Applied Biosystems 7500 Real-Time PCR Instrument system ( Software version 2.0.6), CFX96™ Real-Time PCR Detection system (Software Bio-Rad CFX Maestro version 1.1) | Positive for RdRp with or without E | E gene is for the detection of Sarbecovirus | https://www.fda.gov/media/137935/download
2020-08-05 | Gnomegen LLC | Gnomegen COVID-19-RT-qPCR Detection Kit | Yes | H | N1, N2 | 10 virus copies/reaction | N/A | 3 hours | NP and OP swabs | Applied Biosystems 7500 Fast Dx RealTime PCR Systems (1.4 software) | Positive for both N1 and N2 |  | https://www.fda.gov/media/137895/download
2020-07-05 | Zymo Research Corporation | Quick SARS-CoV-2rRT-PCR Kit | Yes | H | Three regions in N | 83 GEC/mL | N/A | 1.5 hours | Upper respiratory specimens (such as nasal, NP, mid-turbinate or OP swabs), and lower respiratory specimens (such as sputum, tracheal aspirates, and BAL) | Bio-Rad CFX96 Touch Real-Time PCR Detection System ( Bio-Rad CFX Maestro™ 1.1 Version 4.1.2433.1219 software or higher) | Positive for SARS-CoV-2 HEX | The detection of SARS-CoV-2 RNA is performed using three target regions, which are detected in the same fluorescence channel (HEX) | https://www.fda.gov/media/137780/download
2020-06-05 | OPTI Medical Systems, Inc. | OPTI SARS-CoV-2 RT PCR Test | Yes | H | N1, N2 | 0.9 copies/µL (NP swab), 0.7 copies/µL (sputum) | 1800 | 2-3.5 hours | Upper and lower respiratory specimens (such as nasal, NP, OP swabs, sputum, lower respiratory tract aspirates, BAL, and NP wash/aspirate or nasal aspirate) | Applied Biosystems® 7500 FAST (software v2.0.6),Applied Biosystems® QuantStudio 5 ( software v1.5.1), Agilent Mx3005P™ (Software v4.10), Roche LightCycler® 480 (Software v1.5.1), Bio Molecular Systems Mic qPCR (software v2.8.10) | Positive for SARS-CoV-2 FAM | SARS-CoV-2 RNA targets (N1 and N2) are both detected in the FAM channel | https://www.fda.gov/media/137739/download
2020-06-05 | BioMérieux SA | SARS-COV-2 R-GENE | Yes | H | PCR1: N, RdRp PCR2: E | 380 GC/mL | 18000 | 4-5 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and BAL fluid | Applied Biosystems™ 7500 Fast (software version 2.3) or Fast Dx real-time PCR (software version 1.4), QuantStudio 5 Dx (software version 1.0), LightCycler 480 (software version 1.5), CFX96 ( software Maestro 1.1 ) and Rotor‑Gene Q (software 2.3) | PCR1: Positive for one or both of the target genes | E gene is for the detection of Sarbecovirus | https://www.fda.gov/media/137742/download
2020-05-05 | Fast Track Diagnostics Luxembourg S.á.r.l. (a Siemens Healthineers Company) | FTD SARS-CoV-2 | Yes | H | Orf1ab, N | 0.0023 TCID50/mL | N/A | 2-3 hours | Upper respiratory specimens (such as nasal swabs, NP swabs, OP swabs, NP wash/aspirate or nasal aspirate) and BAL | Applied Biosystems® 7500 Real-Time PCR System (ThermoFisher Scientific) and the | Positive for SARS-CoV-2 green channel | Both targets (N gene and ORF1ab) are labeled with the same dye and are detected in the same channel. | https://www.fda.gov/media/137690/download
2020-04-05 | Sansure BioTech Inc. | Novel Coronavirus (2019-nCoV) Nucleic Acid Diagnostic Kit (PCR-Fluorescence Probing) | Yes | H | Orf1ab, N | 200 copies/mL | N/A | 1.15 hours | NP swabs, OP swabs, anterior nasal swabs, midturbinate swabs, nasal washes and nasal aspirates | ABI 7500 Real-Time PCR System | Positive for one or both of the target genes |  | https://www.fda.gov/media/137651/download
04/29/2020 | LabGenomics Co., Ltd. | LabGun COVID-19 RT-PCR Kit | Yes | H | RdRp, E | 20 genomic RNA copies/μl | 1800 | Unavailable | NP or OP, anterior nasal and mid-turbinate nasal swabs, as well as NP wash/aspirate or nasal aspirate specimens and sputum | Applied Biosystems™ 7500 Real-time PCR Instrument system ( software version 2.3) or BioRad CFX96™ Real-time PCR detection system (Bio-Rad CFX Manager 3.1) | Positive for RdRp with or without E |  | https://www.fda.gov/media/137483/download
04/27/2020 | Seasun Biomaterials | U-TOP COVID-19 Detection Kit | Yes | H | 2 regions of Orf1ab gene, N | 1 copy/µL | 600 | 1.5 Hours | OP and NP swab specimens, anterior nasal and mid-turbinate nasal swabs, NP wash/aspirate or nasal aspirate specimens and sputum | CFX 96 real-time PCR detection system ( software CFX manager V3.1) and Applied Biosystems real-time PCR system 7500 (Software 2.0.6) | Positive for one or both of the target genes |  | https://www.fda.gov/media/137425/download
04/23/2020 | SD Biosensor, Inc | STANDARD M nCoV Real-Time Detection Kit | Yes | H | E, RdRp | 0.5 copy/µL (ABI 7500), 0.25 copy/µL for upper respiratory specimens and 0.125 copy/µL for lower respiratory specimens (CFX95), 0.5 copy/µL for upper respiratory specimens and 0.25 copy/µL for lower respiratory specimens (LC480) | N/A | 1 hour | NP, OP, and mid-turbinate nasal swab, and sputum specimens | Roche LightCycler® 480 (software version 1.5.1.62), BioRad CFX96™ Dx System (version 3.1) and Applied Biosystems 7500 Real-Time PCR Instrument System (version 2.0.6) | Positive for RdRp with or without E |  | https://www.fda.gov/media/137302/download
04/22/2020 | altona Diagnostics GmbH | RealStar SARS-CoV-2 RT-PCR Kits U.S. | Yes | H | S, E | 1.0×10-1 PFU/mL | N/A | N/A | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal washes and nasal aspirates | BioRad CFX 96 Touch Real-Time PCR (CFX Manager™ Software v3.1) and CFX96 Touch Deep Well Real-Time PCR Detection Systems (CFX Manager™ Software v3.1) | Positive for S with or without E | E gene is specific for lineage B-betacoronavirus | https://www.fda.gov/media/137252/download
04/21/2020 | Seegene, Inc. | Allplex 2019-nCoV Assay | Yes | H | E, N, RdRp | 4167 copies/mL | N/A | 1 hour 50 mins | NP swab, OP swab, anterior nasal swab, midturbinate nasal swab, nasopharyngeal wash/aspirate, nasal aspirate, BAL and sputum specimens | CFX96TM, CFX96 Touch™ (CFX ManagerTM Software V3.1 or CFX MaestroTM Software V1), Applied Biosystems™ 7500 (software v2.0.5), Applied Biosystems™ 7500 Fast Dx real-time PCR detection systems ( software v1.4.1 (Windows XP)/ 1.5.1 (Windows 7 64-bit)) | Positive for one or both of the target genes (RdRp and N) with or without E |  | https://www.fda.gov/media/137178/download
04/20/2020 | Trax Management Services Inc. | PhoenixDx 2019-CoV | Yes | H | E, RdRp | 100 copies/mL | N/A | 1 hour 15 mins | Nasal swabs, NP swabs, OP swabs and BAL specimens | BIO-RAD CFX96-IVD (CFX Manager Software version 3.0), Qiagen Rotor-Gene Q (software 2.3.5 or higher), and ABI 7500 Fast Real time PCR (software version 2.0.4 or higher) | Positive forRdRP with or without E | E gene is for the detection of Sarbecovirus | https://www.fda.gov/media/137153/download
04/18/2020 | OSANG Healthcare | GeneFinder COVID-19 Plus RealAmp Kit | Yes | H | E, N, RdRP | 0.5 copy/µL | N/A | 3 hours | NP, OP, nasal, and mid-turbinate nasal swab specimens, BAL fluid, and sputum | Applied Biosystems® 7500 Real-Time PCR Instrument (ABI 7500) (software version 2.3), Bio-Rad CFX Real-Time PCR Instrument (software version 1.6) | Positive for one or both of the target genes (RdRp and N) with or without E |  | https://www.fda.gov/media/137116/download
04/17/2020 | Fosun Pharma USA Inc. | Fosun COVID-19 RT-PCR Detection Kit | Yes | H | Orf1ab, E, N | 300 copies/mL | N/A | 1 hour 15 mins | Upper and lower respiratory specimens (such as anterior nasal swabs, mid-turbinate nasal swabs, NP swabs, OP swabs, sputum, lower respiratory tract aspirates, BAL, and NP wash/aspirate or nasal aspirate) | Applied Biosystems™ 7500 Real-Time PCR (software version v1.4 or v1.5) | Positive for two or more of the target genes |  | https://www.fda.gov/media/137120/download
04/16/2020 | KorvaLabs Inc. | Curative-Korva SARS-Cov-2 Assay | No | H | N1, N2 | 200 copies/mL | 18000 | N/A | Nasal swabs, NP swabs, OP swabs and oral fluid specimens | BioRad CFX 96 Touch, BioRad CFX Connect Real-Time PCR systems and Roche LightCycler 480 II Real-Time PCR systems | Positive for SARS-CoV-2 N | The detection of SARS-CoV-2 RNA is performed using two target regions, which are both detected in the same fluorescence channel. | https://www.fda.gov/media/137089/download
04/16/2020 | GenoSensor, LLC | GSTM COVID-19 Real-Time PCR Kit | Yes | H | Orf1ab, E, N | 1 copy/µL | N/A | 1.5 hours | Nasal swabs, NP swabs, OP swabs and mid-turbinate swabs | Applied Biosystems™ 7500 Fast Dx Real-Time PCR (software version 1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/137093/download
04/15/2020 | Maccura Biotechnology (USA) LLC | SARS-CoV-2 Fluorescent PCR Kit | Yes | H | Orf1ab, E, N | 1.0×103 copies/mL (Maccura Nucleic Acid Extraction Kit), 5.0×102 copies/mL (QIAGEN QIAamp Viral RNA Mini Kit) | N/A | 1 hour 15 mins | Upper respiratory specimens (such as oropharyngeal swabs, nasopharyngeal swabs, nasal swabs, and mid-turbinate swabs) | Applied Biosystems™ 7500 Real-Time PCR (software v2.3) | Positive for Orf1ab with or without E/N |  | https://www.fda.gov/media/137026/download
2020-08-04 | DiaCarta, Inc | QuantiVirus™ SARS-CoV-2 Test Kit | Yes | H | Orf1ab, E, N | 100 copies/mL (Orf1ab and N); 200 copies/mL for E (QuantStudio 5); 100 copies/mL (ABI 7500 Fast Dx) | 600 | 2 hours | Nasal swabs, NP swabs, OP swabs and sputum | Thermo Fisher (ABI) QuantStudio 5 (Software v1.4), Thermo Fisher (ABI) 7500 Fast Dx (Software v1.4) | Positive for two or more of the target genes |  | https://www.fda.gov/media/136809/download
2020-07-04 | InBios International, Inc | Smart Detect SARS-CoV-2 rRT-PCR Kit | Yes | H | Orf1ab, E, N | 1.1×103 GE/mL (7500 Fast Dx Real-Time PCR), 8.6×102 GE/mL (CFX96 Touch Real-Time PCR Detection System) | N/A | 1.5 hours | NP swab, anterior nasal swab and midturbinate nasal swab | Applied Biosystems™ 7500 Fast Dx Real-Time PCR (Software version 1.4), CFX96 Touch Real-Time PCR Detection System (Bio-Rad)( CFX Maestro Software) | Positive for two or more of the target genes |  | https://www.fda.gov/media/136786/download
2020-03-04 | Co-Diagnostics, Inc. | LOGIX SMART™ COVID-19 TEST KIT | Yes | H | RdRp | 4.29 copies/µL | N/A | 1-1.5 hours | Lower respiratory tract fluids (e.g. BAL, sputum, tracheal aspirate), and upper respiratory tract fluids (e.g. NP and OP swabs) | CoDx Box (BioMolecular Systems), MIC qPCR Cycler (BioMolecular Systems), Eco 48 (PCRmax) | Positive for RdRp |  | https://www.fda.gov/media/136687/download
2020-03-04 | ScienCell Research Laboratories | ScienCell™ SARS-CoV-2 Coronavirus Real-time RT-PCR (RT-qPCR) Detection Kit | Yes | H | N1, N2 | 100 copies/µL | 540 | 1 hour | Nasal, NP and OP swabs and BAL | Roche LightCycler® 96 Real-Time PCR System (Software 1.01.01.0050) | Positive for one or both of the target genes |  | https://www.fda.gov/media/136691/download
2020-01-04 | Ipsum Diagnostics, LLC | COV-19 IDx assay | No | H | N1 | 8.5 copies/µL | 180000 | 1.5 hours | NP and OP swabs | Applied Biosystems QuantStudio12 Flex (QS12) (software version 1.2.2) | Positive for N1 |  | https://www.fda.gov/media/136621/download
03/27/2020 | Luminex Molecular Diagnostics, Inc. | NxTAG® CoV Extended Panel | Yes | H | Orf1ab, E, N | 5×103 GCE/mL | 5400 | 2.5 hours | NP swab | Luminex platform | Positive for one or more of the target genes |  | https://www.fda.gov/media/136500/download
03/26/2020 | BGI Genomics Co. Ltd | Real-Time Fluorescent RT-PCR Kit for Detecting SARS-2019-nCoV | Yes | H | Orf1ab | 150 copies/mL (throat swab), 100 copies/mL (BAL) | N/A | 2 hours | Throat swabs and BAL fluid | Applied Biosystems™ Real time PCR system 7500 (software v2.0.5) | Positive for Orf1ab |  | https://www.fda.gov/media/136472/download
03/25/2020 | Avellino Lab USA, Inc. | AvellinoCoV2 test | No | H | N1, N3 | 55 copies/µL | 18000 | N/A | NP and OP swabs | Applied Biosystems 7500 Fast Real-Time PCR System (Software version 2.3) | Positive for N1 with or without N3 |  | https://www.fda.gov/media/136453/download
03/24/2020 | PerkinElmer, Inc. | PerkinElmer® New Coronavirus Nucleic Acid Detection Kit | Yes | H | Orf1ab, N | 9.307 copies/ mL (ORF1ab); 30.467 copies/mL (N) | 180 | N/A | NP and OP swabs and anterior nasal swab specimens | Applied BiosystemsTM 7500 Real-Time PCR System (software version 2.3) | Positive for one or both of the target genes |  | https://www.fda.gov/media/136410/download
03/20/2020 | Primerdesign Ltd | Primerdesign Ltd COVID-19 genesig Real-Time PCR assay | Yes | H | Orf1ab | 0.33 copies/µL | N/A | 2 hours | NP and OP swab, BAL | Applied Biosystems® 7500 Real-Time PCR System (software version 2.3), Bio-Rad CFX Connect™ Real-Time PCR Detection System (Maestro™ software version 1.1) and Roche® LightCycler 480 II (software version 1.5) | Positive for Orf1ab |  | https://www.fda.gov/media/136823/download
03/17/2020 | Quest Diagnostics Infectious Disease, Inc. | SARS-CoV-2 RNA Qualitative Real-Time RT-PCR | Yes | H | N1, N3 | 136 copies/mL | N/A | N/A | Upper and lower respiratory specimens (such as NP and OP swabs, sputum, tracheal aspirates, and BAL) | Applied BiosystemsTM 7500 Real-Time PCR System | Positive for both N1 and N3 |  | https://www.fda.gov/media/136231/download
03/17/2020 | Quidel Corp. | Lyra® SARS-CoV-2 Assay | Yes | H | pp1ab | 8×10-1 RNA copies/µL | 6000 | 2.5 hours | NP and OP swabs | Applied Biosystems 7500 Fast Dx (software version 1.4 ), Applied Biosystems 7500 Standard (software version 2.0.6), Roche LightCycler 480 ( software version 1.5.0.39), Qiagen Rotor-Gene Q (software version 2.0.2.4), Bio-Rad CFX96 Touch (software version 3.1), Thermofisher QuantStudio 7 Pro (software version 2.0) | Positive for pp1ab |  | https://www.fda.gov/media/136820/download
03/16/2020 | Laboratory Corporation of America | COVID-19 RT-PCR TEST | No | H | N1, N2, N3 | 6.25 copies/µL (NP), 12.5 copies/µL (BAL) | 18000 | N/A | Upper and lower respiratory specimens (such as NP and OP swabs, sputum, lower respiratory tract aspirates, BAL, and NP wash/aspirate or nasal aspirate) | Applied Biosystems QuantStudio7 Flex (QS7) (software version 1.3) | Positive for all three target genes |  | https://www.fda.gov/media/136151/download
03/13/2020 | Thermo Fisher Scientific, Inc. | TaqPath™ COVID-19 Combo Kit | Yes | H | Orf1ab, S, N | 10 GCE/ reaction | 180000 | 3 hours | Upper respiratory specimens (such as NP, OP, nasal, and mid-turbinate swabs, and NP aspirate) and BAL specimens | Applied Biosystems™ 7500 Fast Dx Real-Time PCR ( Software v1.4.1), Applied Biosystems™ 7500 Fast Real‑Time PCR Instrument (Software v1.5.1 or v2.3), Applied Biosystems™ 7500 Real‑Time PCR Instrument (Software v2.3), Applied Biosystems™ QuantStudio™ 5 Real‑Time PCR (Software v1.5.1), Applied Biosystems™ QuantStudio™ 7 Flex Real-Time PCR (Software v1.3) | Positive for two or more of the target genes |  | https://www.fda.gov/media/136112/download
02/29/2020 | Wadsworth Center, NYSDOH | New York SARS-CoV-2 Real-time Reverse Transcriptase (RT)-PCR Diagnostic Panel | Yes | H | N1, N2 | 25 GC/ reaction | N/A | N/A | NP, OP swabs and sputa | Applied Biosystems® 7500 Fast Dx Real-Time PCR (Software version 1.4) | Positive for both N1 and N2 |  | https://www.fda.gov/media/135847/download
2020-04-02 | Centers for Disease Control and Prevention's (CDC) | CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel | Yes | H | N1, N2 | 100 RNA copies/µL (Qiagen DSP); 100.5 RNA copies/µL (Qiagen EZ1) | 18000 | 3 hours | Upper and lower respiratory specimens (such as NP and OP swabs, sputum, lower respiratory tract aspirates, BAL, and NP wash/aspirate or nasal aspirate) | Applied Biosystems 7500 Fast Dx Real-Time PCR Instrument (Software v1.4) | Positive for both N1 and N2 |  | https://www.fda.gov/media/134922/download
Real-time RT-PCR for Use on Automated Sample-to-result System
2020-02-10 | BioFire Diagnostics, LLC | BioFire Respiratory Panel 2.1-EZ (RP2.1-EZ) | Yes | H, M, W | S, M | 5×102 copies/mL | N/A | N/A | NP swab | BioFire® FilmArray® System instrument module | Positive for one or both of the target genes |  | https://www.fda.gov/media/142696/download
09/24/2020 | Cepheid | Xpert Xpress SARS-CoV-2/Flu/RSV | Yes | H, M, W | E, N2 | 131 copies/mL | N/A | 2-4 hours | NP swab, nasal swab or nasal wash/ aspirate specimens | GeneXpert Dx, GeneXpert Infinity and/or GeneXpert Xpress systems | Positive for SARS-CoV-2 channel |  | https://www.fda.gov/media/142437/download
2020-04-09 | Roche Molecular Systems, Inc. | cobas SARS-CoV-2 & Influenza A/B Nucleic Acid Test for use on the cobas Liat System | Yes | H, M, W | Orf1a/b, E | 12 copies/mL | N/A | 2-4 hours | NP and nasal swabs | cobas® Liat® System | Positive for Orf1a/b with or without E |  | https://www.fda.gov/media/142193/download
2020-03-09 | Roche Molecular Systems, Inc. | cobas SARS-CoV-2 & Influenza A/B | Yes | H, M | Orf1a/b, E | 0.12 TCID50/mL | 1800 | 2-4 hours | Nasal and NP swab specimens | cobas® 6800 System and cobas® 8800 System | Positive for Orf1a/b with or without E | E gene is specific for pan-Sarbecovirus | https://www.fda.gov/media/141887/download
08/31/2020 | BayCare Laboratories, LLC | BayCare SARS-CoV-2 RT PCR Assay | No | H | ORF1, E | 0.007 TCID50/mL | N/A | 2-4 hours | NP swabs | cobas 6800/8800 Systems with cobas 6800/8800 Systems Software v1.2 (ASAP v10.1.0) and v1.3 (ASAP v11.1.0) | Positive for one or both of the target genes |  | https://www.fda.gov/media/141769/download
06/13/2020 | Kaiser Permanente Mid-Atlantic States | KPMAS COVID-19 Test | No | H | Orf1ab, E | 0.007 TCID50/mL (Orf1ab); 0.004 TCID50/mL € | N/A | 3 hours | Nasal swab | cobas® 6800/8800 Systems | Positive for both Orf1ab and E |  | https://www.fda.gov/media/139067/download
07/31/2020 | University of California San Diego Health | UCSD RC SARS-CoV-2 Assay | No | H | Orf1ab, E | 0.007 TCID50/mL (Orf1ab); 0.004 TCID50/mL € | N/A | 2 hours | Nasal swab specimens and upper respiratory specimens (NP, mid-turbinate, anterior nares or OP swab specimens) | cobas 6800/8800 Systems with cobas 6800/8800 Systems Software v1.2 (ASAP v10.1.0) and v1.3 (ASAP v11.1.0) | Positive for one or both of the target genes |  | https://www.fda.gov/media/140712/download
07/15/2020 | Quest Diagnostics Infectious Disease, Inc. | Quest Diagnostics PF SARS-CoV-2 Assay | No | H | 2 regions of Orf1ab | 1×10-2 TCID50/mL | N/A | N/A | Nasal swab | Panther Fusion® System | Positive for SARS-CoV-2 ROX channel |  | https://www.fda.gov/media/140229/download
07/15/2020 | Quest Diagnostics Infectious Disease, Inc. | Quest Diagnostics RC SARS-CoV-2 Assay | No | H | Orf1ab, E | 0.007 TCID50/mL (Orf1ab); 0.004 TCID50/mL € | N/A | Unavailable | Nasal swab | cobas® 6800/8800 Systems | Positive for both Orf1ab and E |  | https://www.fda.gov/media/140234/download
06/17/2020 | Omnipathology Solutions Medical Corporation | Omni COVID-19 Assay by RT-PCR | No | H | N1, N2 | 1.23 copies/µL | N/A | 45 minutes | NP swab, OP swab, anterior nasal swab, mid-turbinate nasal swab, NP wash/aspirate, nasal aspirate, and BAL specimens | BD MAX System | Positive for SARS-CoV-2 channel | The detection of SARS-CoV-2 RNA is performed using two target regions, which are both detected in the same fluorescence channel | https://www.fda.gov/media/139292/download
2020-11-05 | Abbott Molecular Inc. | Alinity m SARS-CoV-2 assay | Yes | H, M | N, RdRp | 100 virus copies/mL | N/A | <2 hours | Nasal swabs, NP and OP swabs and BAL | Alinity m System | Positive for SARS-CoV-2 channel | The two SARS-CoV-2-specific probes are labeled with the same fluorophore | https://www.fda.gov/media/137979/download
2020-01-05 | BioFire Diagnostics, LLC | BioFire Respiratory Panel 2.1 (RP2.1) | Yes | H, M | S, M | 5.0×102 copies/mL | N/A | 45 minutes | NP swab | BioFire® FilmArray® System instrument module | Positive for one or both of the target genes |  | https://www.fda.gov/media/137583/download
2020-08-04 | Becton, Dickinson & Company (BD) | BD SARS-CoV-2Reagents for BD MAX System | Yes | H, M | N1, N2 | 40 GE/mL | 5400 | N/A | NP, nasal, mid-turbinate, and OP swab specimens, NP wash/aspirate or nasal aspirates | BD MAX System | Positive for one or both of the target genes |  | https://www.fda.gov/media/136816/download
2020-03-04 | Luminex Corporation | ARIES® SARS-CoV-2 Assay | Yes | H, M | Orf1ab, N | 3.33×102 GCE/mL | 180000 | 4 hours | NP swab | ARIES instrument | Positive for one or both of the target genes |  | https://www.fda.gov/media/136693/download
2020-02-04 | Becton, Dickinson & Company (BD) | BioGX SARS-CoV-2 Reagents for BD MAX™ System | Yes | H, M | N1, N2 | 40 GE/mL | 1800 | 3 hours | NP, nasal, mid-turbinate, and OP swab specimens, nasopharyngeal wash/aspirate or nasal aspirates | BD MAX System | Positive for one or both of the target genes |  | https://www.fda.gov/media/136653/download
03/30/2020 | QIAGEN GmbH | QIAstat-Dx® Respiratory SARS-CoV-2 Panel | Yes | H, M | Orf1b, E | 500 copies/mL | 180000 | N/A | NP swab | QIAstat-Dx Analyzer 1.0 | Positive for SARS-CoV-2 channel | The two SARS-CoV-2-specific probes are labeled with the same fluorophore | https://www.fda.gov/media/136571/download
03/30/2020 | NeuMoDx Molecular, Inc. | NeuMoDx™ SARS-CoV-2 Assay | Yes | H, M | N, Nsp2 | 150 copies/mL | 5400 | N/A | Nasal, NP and OP swabs, and BAL specimens | NeuMoDx Molecular System | Positive for one or both of the target genes |  | https://www.fda.gov/media/136565/download
03/23/2020 | BioFire Defense, LLC | BioFire® COVID-19 Test | Yes | H, M | 2 regions of Orf1ab, Orf8 | 3.3×102 GC/mL | 5400 | 45 minutes | NP swab | FilmArray® 2.0 or FilmArray® Torch systems | Positive for two or more of the target genes |  | https://www.fda.gov/media/136353/download
03/20/2020 | Cepheid | Xpert® Xpress SARS-CoV-2 | Yes | H, M, W | E, N2 | 0.02 PFU/mL | 5400 | 45 minutes | Upper respiratory specimens (such as NP, OP, nasal, or mid-turbinate swab and/or nasal wash/ aspirate) | GeneXpert Dx and GeneXpert Infinity systems | Positive for N2 wth or without E |  | https://www.fda.gov/media/136314/download
03/19/2020 | DiaSorin Molecular LLC | Simplexa™ COVID-19 Direct | Yes | H, M | Orf1ab, S | 500 copies/mL | 6000 | 1 hour | NP swabs, nasal swabs, nasal wash/aspirate or BAL specimens | LIAISON® MDX | Positive for one or both of the target genes |  | https://www.fda.gov/media/136286/download
03/18/2020 | Abbott Diagnostics Scarborough, Inc. | Abbott RealTime SARS-CoV-2 | Yes | H | N, RdRp | 100 virus copies/mL | 5400 | 24 hours (high volume throughout) | Nasal, NP and OP swabs, and BAL | Abbott m2000 System | Positive for SARS-CoV-2 channel | The two SARS-CoV-2-specific probes are labeled with the same fluorophore | https://www.fda.gov/media/136258/download
03/16/2020 | Hologic, Inc. | Panther Fusion SARS-CoV-2 Assay | Yes | H | 2 regions of Orf1ab | 1×10-2 TCID50/mL | 600 | 3.5 hours | Upper respiratory specimens (such as NP, nasal, mid-turbinate, and OP swab specimens, NPwash/aspirates or nasal wash), and lower respiratory tract (LRT) specimens (such as BAL) | Panther Fusion® System | Positive for SARS-CoV-2 ROX channel | The two SARS-CoV-2-specific probes are labeled with the same fluorophore ROX | https://www.fda.gov/media/136156/download
2020-12-03 | Roche Molecular Systems, Inc. (RMS) | cobas® SARS-CoV-2 | Yes | H, M | Orf1ab, E | 0.007 TCID50/mL (Orf1ab); 0.004 TCID50/mL € | 1800 | 3 hours | Nasal, NP and OP swabs | cobas® 6800/8800 Systems | Positive for both Orf1ab and E |  | https://www.fda.gov/media/136049/download
Endpoint RT-PCR
Lateral Flow Detection
03/23/2020 | Mesa Biotech Inc. | Accula SARS-Cov-2 Test | Yes | H, M, W | N | 200 copies/reaction | N/A | 30 minutes | Nasal or nasal mid-turbinate swab specimens | Accula™ Dock or Silaris™ Dock | Positive for N |  | https://www.fda.gov/media/136355/download
Enzyme-based Colorimetric Detection
09/16/2020 | Visby Medical, Inc. | Visby Medical COVID-19 | Yes | H, M | N | 1112 GC/mL | N/A | 2-4 hours | NP, nasal, or mid-turbinate swabs | Instrument-free | Positive for N |  | https://www.fda.gov/media/142228/download
04/29/2020 | Rheonix, Inc. | Rheonix COVID-19 MDx Assay | Yes | H | N1 | 625 GE/mL | 1800 | <5 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal washes, nasal aspirates and BAL fluid | Rheonix Encompass MDx® Workstation | Positive for N1 |  | https://www.fda.gov/media/137489/download
Fluorescence Detection
09/30/2020 | Alimetrix, Inc. | Alimetrix SARS-CoV-2 RT-PCR Assay | No | H | Orf1ab, N1, N2 | 250 copies/mL (Zymo Research Quick-DNA/RNA Viral MagBead Extraction); 1000 copies/mL (Qiagen QIAamp 96 Virus QIAcube HT Kit) | N/A | 2-4 hours | NP, OP, anterior nasal and mid-turbinate nasal swab specimens, as well as NP washes/aspirates or nasal aspirates and BAL specimens | Applied Biosystems GeneAmp 9700 Thermal Cycler with software version 3.12, Sensovation Sensopot microarray scanner with ArrayReader software version 3.1.0.15245 | Positive for one or more of the target genes |  | https://www.fda.gov/media/142592/download
08/28/2020 | DxTerity Diagnostics, Inc. | DxTerity SARS-CoV-2 RT PCR CE Test | No | H | Orf1ab, E, N | 50 copies/mL | N/A | N/A | Saliva specimens | VeritiDx PCR Thermal Cycler, ThermoFisher ABI 3500xL Dx Genetic Analyzer (K191030) | Positive for two or more of the target genes |  | https://www.fda.gov/media/141669/download
08/25/2020 | QDx Pathology Services | QDX SARS-CoV-2 Assay | No | H | N1, N2 | 1000 copies/mL (Applied Biosystems 7500 Fast and the Applied Biosystems Quant Studio 7 systems), 250 copies/mL (Applied Biosystems Quant Studio 12K system) | N/A | N/A | Nasal swab | Applied Biosystems™ QuantStudio™ 7 Flex (Software Version 1.3), Applied Biosystems™ 7500 Fast (Software Version 2.3), Applied Biosystems™ QuantStudio™ 12K Flex (Software Version 1.3) | Positive for N1 and N2 |  | https://www.fda.gov/media/141545/download
06/25/2020 | PlexBio Co., Ltd. | IntelliPlex SARS-CoV-2 Detection Kit | Yes | H | E, N, RdRP | 140 copies/mL | N/A | N/A | NP, OP, anterior nasal and mid-turbinate nasal swabs, NP wash/aspirate or nasal aspirates, and BAL specimens | IntelliPlex 1000 πCode Processor and PlexBio 100 Fluorescent Analyzer with DeXipher software | Positive for one or more of the target genes |  | https://www.fda.gov/media/139527/download
06/15/2020 | Applied BioCode, Inc. | BioCode SARS-CoV-2 Assay | Yes | H | 2 regions in N gene | 1.72×10-2 TCID50/mL | 5400 | N/A | Upper respiratory specimens (such as NP swabs, OP swabs, and nasal swabs) or BAL | BioCode® MDx-3000 | Positive for one or both of the target genes |  | https://www.fda.gov/media/139049/download
2020-09-06 | ChromaCode Inc. | HDPCR SARS-CoV-2 Assay | Yes | H | N1, N2 | 1000 copies/mL (Applied Biosystems 7500 Fast and the Applied Biosystems Quant Studio 7 systems), 250 copies/mL (Applied Biosystems Quant Studio 12K system) | 5400 | <1.5 hours | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs as well as nasal aspirate, nasal wash, and BAL specimens | Applied Biosystems™ QuantStudio™ 7 Flex (Software Version 1.3), Applied Biosystems™ 7500 Fast (Software Version 2.3), Applied Biosystems™ QuantStudio™ 12K Flex (Software Version 1.3) | Positive for N1 and N2 |  | https://www.fda.gov/media/138786/download
Electrochemical detection
2020-07-10 | GenMark Diagnostics, Inc. | ePlex Respiratory Pathogen Panel 2 | Yes | H, M | NR | 250 GC/mL | N/A | 2-4 hours | NP swab | ePlex | Positive for the target gene |  | https://www.fda.gov/media/142905/download
03/19/2020 | GenMark Diagnostics, Inc. | ePlex®SARS-CoV-2 Test | Yes | H, M | NR | 750 GC/mL | N/A | <2 hours | NP swab | ePlex | Positive for the target gene |  | https://www.fda.gov/media/136282/download
Magnetic Resonance Detection
08/31/2020 | T2 Biosystems, Inc. | T2SARS-CoV-2 Panel | Yes | H, M | NR | 2000 GE/mL | N/A | N/A | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swab specimens) and BAL specimens | T2Dx Instrument (T2Dx) | Positive for the target gene |  | https://www.fda.gov/media/141755/download
MALDI-TOF Detection
10/26/2020 | Agena Bioscience, Inc. | MassARRAY SARS-CoV-2 Panel | Yes | H | Orf1ab, N1, N2, N3, ORF1 | 2.5 copies/μl | N/A | N/A | Upper respiratory specimens (NP swab, OP swab, nasal and mid-turbinate swabs, and nasal and NP aspirate) and BAL | MassARRAY System with Typer Software | Positive for two or more of the target genes |  | https://www.fda.gov/media/143334/download
09/29/2020 | National Jewish Health | SARS-CoV-2 MassArray Test | No | H | Orf1ab, N1, N2, N3, ORF1 | 0.69 copies/µL | 5400 | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP, and OP swabs and nasal/NP aspirate and wash specimens) and BAL specimens | Thermo Fisher Veriti thermal cycler, MassARRAY System with SpectroACQUIRE software v4.3.145 and v.5.0 | Positive for two or more of the target genes |  | https://www.fda.gov/media/142548/download
2020-03-08 | Ethos Laboratories | Ethos Laboratories SARS-CoV-2 MALDI-TOF Assay | No | H | Orf1ab, N1, N2, N3, ORF1 | 1 TCID50/mL | 5400 | 2-4 hours | Upper respiratory specimens (such as nasal, mid-turbinate, NP and OP swab specimens) and BAL specimens | Veriti 96 Thermal Cycler with Thermal Cycler Fleet Control Software Version 1.7.1 (Applied Biosystems, California, USA), MassARRAY System with Typer Software, Version 5.0 | Positive for two or more of the target genes |  | https://www.fda.gov/media/140780/download
RT-Digital PCR
06/25/2020 | PreciGenome LLC | FastPlex Triplex SARS-CoV-2 detection kit (RT-Digital PCR) | Yes | H | Orf1ab, N | 571.4 copies/mL | 180000 | ~1-2 hours | OP swabs | DropX-2000 Sample Prep Station and Drop-X 2000 DScanner | Positive for one or both of the target genes |  | https://www.fda.gov/media/139523/download
2020-06-04 | Gnomegen LLC | Gnomegen COVID-19 RT-Digital PCR Detection Kit | Yes | H | N1, N2 | 8 GC/reaction | N/A | N/A | Nasal, NP and OP swab | QuantStudio™ 3D Digital PCR System | Positive for SARS-CoV-2 FAM channel | The two SARS-CoV-2-specific probes are labeled with the same fluorophore FAM | https://www.fda.gov/media/136738/download
2020-01-05 | Bio-Rad Laboratories, Inc | Bio-Rad SARS-CoV-2 ddPCR Test | Yes | H | N1, N2 | 150 copies/mL | 600 | 45 minutes | NP, anterior nasal and mid-turbinate nasal swab specimens as well as NP wash/aspirate and nasal aspirate specimens | QX200/QXDx Automated Droplet Generator, Bio-Rad C1000 Touch or S1000 Thermocyclers and QX200/QXDx Droplet Reader | Positive for one or both of the target genes |  | https://www.fda.gov/media/137579/download
qSTAR
10/14/2020 | LumiraDx UK Ltd. | LumiraDx SARS-CoV-2 RNA STAR Complete | Yes | H | Orf1a | 7500 copies/mL | N/A | 12 minutes | Upper respiratory specimens (such as nasal, mid-turbinate, NP and OP swabs) | Roche LightCycler 480 II (software version SW 1.5.1), Applied Biosystems 7500 Fast Dx (software version 1.4.1), Applied Biosystems QuantStudio 5 (software version 1.5.1), Agilent AriaMx (software version 1.71), or the Agilent Stratagene Mx3005P (software version 4.10) RT-PCR Instruments | Positive for Orf1a |  | https://www.fda.gov/media/143062/download
2020-11-08 | LumiraDx UK Ltd. | LumiraDx SARS-CoV-2 RNA STAR | Yes | H | Orf1a | 500 copies/mL | N/A | N/A | Upper respiratory specimens (such as nasal, midturbinate, NP and OP swabs) and BAL specimens | Roche LightCycler 480 II (software version SW 1.5.1), Applied Biosystems 7500 Fast Dx (software version 1.4.1), Applied Biosystems QuantStudio 5 (software version 1.5.1), Agilent AriaMx (software version 1.71), or the Agilent Stratagene Mx3005P (software version 4.10) RT-PCR Instruments | Positive for Orf1a |  | https://www.fda.gov/media/141057/download
Isothermal Nucleic Acid Amplification
2020-10-06 | Cue Health Inc. | Cue COVID-19 Test | Yes | H, M, W | N | 1.3 GC/µL | N/A | 25 minutes | Direct nasal swab or nasal specimens | Cue Health Monitoring System, Apple® iPhone® 8+ or later mobile smart device with iOS 13 or later with Bluetooth and Wi- Fi | Postiitve for N |  | https://www.fda.gov/media/138826/download
Fluorescence Detection
2020-05-10 | SEASUN BIOMATERIALS, Inc. | AQ-TOP COVID-19 Rapid Detection Kit PLUS | Yes | H | Orf1ab, N | 1 copy/µL | N/A | N/A | Upper respiratory specimens including OP and NP swabs, anterior nasal and mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates as well as BAL specimens | CFX 96 real-time PCR detection system with software CFX manager V3.1, Applied Biosystems real-time PCR system 7500 with Software 2.0.6 | Positive for one or both target geens |  | https://www.fda.gov/media/142800/download
08/13/2020 | Pro-Lab Diagnostics | Pro-AmpRT SARS-CoV-2 Test | No | H | Orf1ab | 125 GE/swab | N/A | 2.5 hours | Upper respiratory specimens (such as nasal, NP and OP swab specimens) | Genie HT instrument (Software version v4.00rc2) | Positive for Orf1ab |  | https://www.fda.gov/media/141149/download
05/21/2020 | Seasun Biomaterials, Inc. | AQ-TOP COVID-19 Rapid Detection Kit | Yes | H | Orf1ab | 7 copies/ µL | 6000 | 15-30 minutes | Upper and lower respiratory specimens including OP and NP swab specimens, anterior nasal and mid-turbinate nasal swabs, NP wash/aspirate or nasal aspirate specimens, BAL and sputum | CFX 96 real-time PCR detection system (Bio-Rad) (Software CFX manager V3.1) and ABI 7500 realtime PCR platforms (Software 2.0.6) | Positive for Orf1ab |  | https://www.fda.gov/media/138307/download
Colorimetric Detection
2020-01-09 | Detectachem Inc. | MobileDetect Bio BCC19 (MD-Bio BCC19) Test Kit | Yes | H, M | E, N | 75 copies/µL | N/A | 4-6 hours | NP, OP, mid-turbinate (MT) and anterior nares (nasal) swab specimens | MD-Bio BCC19 Heater | positive for one or both of the target genes |  | https://www.fda.gov/media/141791/download
05/18/2020 | Color Genomics, Inc. | Color Genomics SARS-CoV-2 RT-LAMP Diagnostic Assay | No | H | Orf1ab, E, N | 0.75 copies/µL | 18000 | N/A | NP swabs, OP swabs, anterior nares swabs, mid-turbinate nasal swabs, NP washes/aspirates or nasal aspirates, as well as BAL specimens | Biotek NEO microplate | Positive for all the three target genes |  | https://www.fda.gov/media/138249/download
CRISPR-based Detection
08/31/2020 | Mammoth Biosciences, Inc. | SARS-CoV-2 DETECTR Reagent Kit | Yes | H | N | 20 copies/µL | N/A | 2-4 hours | Upper respiratory specimens (NP swabs, OP swabs, mid-turbinate nasal swabs, anterior nasal swabs, NP wash/aspirate or nasal aspirate) | ABI 7500 Fast Dx Real-Time PCR system | Positive for N |  | https://www.fda.gov/media/141765/download
2020-09-07 | UCSF Health Clinical Laboratories, UCSF Clinical Labs at China Basin | SARS-CoV-2 RNA DETECTR Assay | No | H | N | 20 copies/µL | N/A | N/A | Upper respiratory specimens (such as nasal swabs, NP swabs, OP swabs, NP wash/aspirate or nasal aspirate) and BAL specimens | ABI 7500 Fast Dx instrument (Applied Biosciences) (Software Version 1.4.1) | Postive for N |  | https://www.fda.gov/media/139937/download
2020-06-05 | Sherlock BioSciences, Inc. | Sherlock CRISPR SARS-CoV-2 Kit | Yes | H | Orf1ab, N | 6.75 copies/µL (Orf1ab); 1.35 copies/µL (N) | 6000 | 1 hour | Upper respiratory specimens (such as nasal swabs, NP swabs, OP swabs, NP wash/aspirate or nasal aspirate) and BAL specimens | BioTek Plate Reader running Gen5 3.08 software | Positive for one or both of the target genes |  | https://www.fda.gov/media/137746/download
NEAR
03/27/2020 | Abbott Diagnostics Scarborough, Inc. | ID NOW COVID-19 | Yes | H, M, W | RdRp | 125 GE/mL | 300000 | 13 minutes | Nasal, NP or throat swabs | ID NOW Instrument | Positive for RdRp |  | https://www.fda.gov/media/136525/download
OMEGA
2020-10-04 | Atila BioSystems, Inc. | iAMP COVID-19 Detection Kit | Yes | H | Orf1ab, N | 10 copies/μL | N/A | N/A | Nasal, NP and OP swabs | CFX96 Real-Time System | Positive for SARS-CoV-2 FAM channel |  | https://www.fda.gov/media/136870/download
TMA
2020-03-08 | Poplar Healthcare | Poplar SARS-CoV-2 TMA Pooling assay | No | H | 2 regions of Orf1ab | 83 copies/mL | N/A | N/A | Nasal, NP and OP swab specimens | Panther Fusion® system | Positive for SARS-CoV-2 channel |  | https://www.fda.gov/media/140792/download
07/15/2020 | Quest Diagnostics Infectious Disease, Inc. | Quest Diagnostics HA SARS-CoV-2 Assay | No | H | 2 regions of Orf1ab | 83 copies/mL | N/A | N/A | Nasal swab | Panther Fusion® system | Positive for SARS-CoV-2 channel |  | https://www.fda.gov/media/140239/download
05/28/2020 | PrivaPath Diagnostics, Inc. | LetsGetChecked Coronavirus (COVID-19) Test | No | H | 2 regions of Orf1ab | 83 copies/mL | 720 | N/A | Nasal swab | Panther Fusion® system | Positive for SARS-CoV-2 channel |  | https://www.fda.gov/media/138406/download
05/14/2020 | Hologic, Inc. | Aptima SARS-CoV-2 assay | Yes | H | 2 regions of Orf1ab | 83 copies/mL | 600 | 2.4 hours | NP, nasal, mid-turbinate and OP swab specimens, nasopharyngeal wash/aspirate or nasal aspirates | Panther Fusion® system | Positive for SARS-CoV-2 channel |  | https://www.fda.gov/media/138096/download
Sequencing
Sanger Sequencing
2020-04-09 | BillionToOne, Inc. | qSanger-COVID-19 Assay | Yes | H | N | 3200 copies/mL | N/A | 2-4 hours | Upper respiratory swab specimens (such as nasal swab, mid-turbinate swab, NP swab, and OP swab specimens) | Thermal cycler (e.g. Applied Biosystems Veriti Thermal Cycler) and Sanger sequencing instrument (e.g. Applied Biosystems 3730xl DNA Analyzer) | Positive for N |  | https://www.fda.gov/media/141935/download
NGS
2020-06-10 | University of California, Los Angeles (UCLA | UCLA SwabSeq COVID-19 Diagnostic Platform | No | H | S2 | 250 GCE/mL | N/A | N/A | Upper respiratory specimens (NP swabs, OP swab, midturbinate nasal swabs and anterior nasal swabs) | Veriti 384-Well Thermal Cycler, Illumina MiSeq sequencer | S2/S2 spike ratio > 0.003 |  | https://www.fda.gov/media/142805/download
09/23/2020 | Clear Labs, Inc. | Clear Dx SARS-CoV-2 Test | Yes | H | 21 target genes | 2000 copies/mL | N/A | 2-4 hours | NP swab, OP swab, anterior nasal swab, midturbinate swab, NP wash/aspirate, nasal aspirate, and BAL specimens | Clear Dx™ system including Hamilton STAR robotic platform and Oxford Nanopore GridION Sequencer | Alogorithm |  | https://www.fda.gov/media/142418/download
08/21/2020 | Guardant Health, Inc. | Guardant-19 | No | H | N1 | 125 copies/mL | N/A | 2-4 hours | NP swabs, OP swabs, nasal swabs, anterior nasal swabs, mid-turbinate nasal swabs, NP wash/aspirates, nasal aspirates, and nasal washes | Illumina NextSeq 500 and NextSeq 550 Sequencing Systems | Positive for N1 |  | https://www.fda.gov/media/141487/download
2020-06-08 | Helix OpCo LLC (dba Helix) | Helix COVID-19 NGS Test | No | H | S | 125 GCE/mL | N/A | 2-4 hours | Upper respiratory specimens (NP swabs, OP swab, midturbinate nasal swabs and anterior nasal swabs) | NovaSeq 6000 Sequencing System | Positive for S |  | https://www.fda.gov/media/140917/download
2020-09-06 | Illumina, Inc. | Illumina COVIDSeq Test | Yes | H | 98 targets | 1000 copies/mL | 5400 | N/A | NP swabs, OP swabs, anterior nasal swabs, mid-turbinate nasal swabs, NP wash/aspirates, nasal aspirates, and BAL specimens | NovaSeq 6000 Sequencing System, NextSeq 500 Sequencing System, NextSeq 550 Sequencing System, or NextSeq 550Dx Instrument | At least 90 SARS-CoV-2 virus targets detected |  | https://www.fda.gov/media/138776/download
NDU, NAAT Detectable Units; GE, genomic equivalents; GC, genomic copies; GCE, genomic copy equivalents; PFU, plaque forming units; NP, nasopharyngeal; OP, oropharyngeal; N/A, not available.
*https://www.fda.gov/medical-devices/coronavirus-covid-19-and-medical-devices/sars-cov-2-reference-panel-comparative-data (Accessed on 7 Nov 2020)
**https://ifh.rutgers.edu/covid-test/ (Accessed on 7 Nov 2020)